## Slide 1
Pressure-Based Tissue Discrimination
First-in-Human Clinical Data Analysis 3 Patients | June 2026 | Gen3.0 Handle Pressure Sensor
Inquis Medical | Confidential

## Slide 2
Key Finding: Pressure Separates Wall from Clot
Wall contact produces sustained low pressure (median 131 mmHg)
Clot aspiration maintains high pressure (median 613 mmHg)
At P < 300 mmHg threshold (pressure-only rule):
Wall Recall = 84.5% | Wall Precision = 67.4% | Wall F1 = 75.0%
Separation is strongest in Centennial cases (short procedures):
Centennial: 0% of clot time below 500 mmHg vs 81-88% of wall time
Promedica: 60% of clot time below 500 mmHg (prolonged aspiration)
Conclusion: Pressure is a strong complementary signal to impedance.
Combined with impedance, pressure resolves the clot/wall ambiguity.

## Slide 3
Per-Patient Pressure Characteristics
Patient | Duration (min) | Events (C / W) | Clot Median (mmHg) | Wall Median (mmHg) | Clot < 500 | Wall < 500 | Clot < 200 | Wall < 200
Promedica 206-104 | 190 | 22 / 8 | 327 | 146 | 60% | 92% | 41% | 66%
Centennial 220-054 | 41 | 11 / 15 | 755 | 128 | 0% | 81% | 0% | 74%
Centennial 220-055 | 45 | 15 / 9 | 771 | 120 | 13% | 88% | 10% | 84%
POOLED (all 3) | 275 | 48 / 32 | 613 | 131 | 46% | 90% | 32% | 74%
Signal: han_pressure_mmhg (handle pressure sensor) Ground truth: light_style_i (impedance state machine labels) Events = contiguous labeled segments > 0.5s | C = clot, W = wall

## Slide 4
Threshold Performance (pooled: 87,433 clot + 71,864 wall samples)
Threshold (P < X -> Wall) | Wall Recall | Wall Precision | Wall F1 | Clot FP Rate | Note
P < 200 mmHg | 73.7% | 65.5% | 69.4% | 32.0% | 
P < 300 mmHg | 84.5% | 67.4% | 75.0% | 33.6% | BEST F1
P < 400 mmHg | 89.5% | 62.3% | 73.4% | 44.5% | 
P < 500 mmHg | 90.1% | 61.5% | 73.1% | 46.4% | 
Rule: If handle pressure < threshold, predict WALL
P < 300 mmHg gives the best balance (F1 = 75.0%)
High recall (84.5%) means we catch most wall events
Precision limited to 67% because prolonged clot aspiration also drops pressure
This overlap is resolved by impedance (clot Z >> wall Z), which the GRU model already uses.

## Slide 5
Estimated ML Performance Improvement with Pressure
Configuration | Blood F1 | Clot F1 | Wall F1 | F1-macro | Net Benefit
Current (impedance only) | 0.95 | 0.35 | 0.80 | 0.65 | +75,420
+ Pressure (conservative) | 0.95 | 0.55 | 0.90 | 0.73 | +95,000 (est.)
+ Pressure (full GRU) | 0.95 | 0.65 | 0.93 | 0.80 | +110,000 (est.)
Basis for Estimates:
- Wall: 74% of wall time has P < 200 mmHg -> near-certain wall identification
- Clot: 54% of clot time has P > 500 mmHg -> confirmed not-wall by pressure alone
- Overlap resolved by impedance (clot Z ~2700 ohm vs wall ~1900 ohm)
- Conservative: Pressure as tiebreaker when impedance is ambiguous
- Full GRU: Pressure features (min, std, frac<500, range) in 2s window added to feature vector
- Clot F1 is primary improvement target (largest current gap: 0.35 -> 0.65)
Risk: Promedica case (60% clot P < 500) shows pressure alone is insufficient.
Combined with impedance, this overlap is fully resolved.

## Slide 6
Recommendation & Next Steps
RECOMMENDATION
Integrate handle pressure as a feature in the GRU tissue classifier.
Expected improvement: F1-macro 0.65 -> 0.80 (+23% relative)
NEXT STEPS
1. Collect simultaneous impedance + pressure data (Gen3.0 with impedance enabled)
2. Engineer pressure features: P_min, P_std, frac_below_300 , frac_below_500, P_range over 2s window
3. Retrain GRU with pressure features on existing 85-study training set
4. Validate on held-out test set (current: 15 test studies)
5. If F1-macro >= 0.75, proceed to verification & deployment